## Slide 1: Detector Simulations Using Geant4
9th International Geant4 Tutorial in Korea 2022
Young-su Kim
Korea Atomic Energy Research Institute
Notes: Hello. I am Young-su Kim in Korea Atomic Energy Research Institute, KAERI.
Today I would like to share my experience on detector simulations using Geant4.
First of all I would like to apologize to present this in Korean. Because I think it will not be well delivered if I use English.

## Slide 2: Contents
Scintillator & Imaging System Modeling
Gamma-ray Spectrum Generator
Closing
Notes: First item is scintillation detector and imaging system modeling and the second one is gamma-ray spectrum generator.

## Slide 3: Scintillator & Imaging System Modeling

## Slide 4: Principle of Compton Camera
Compton camera — the eye to see radiation
-emitting source
Scatter
Absorber
Notes: The imaging system, I participated to develop one, was Compton camera.
Compton camera is so-called the eye to see radiation, which is able to image the location of gamma-ray emitting radio isotopes.
By using the formula to calculate Compton scattering angle, shown here, and by using the energy and position information of each component detector,
we can draw a cone that cross the point of the source gamma emitted.
By overlapping these cones over and over, the source location can be imaged.

## Slide 5: Application of Compton Camera (1/2)
- Localization of primary source term / clean-up verification
- Locating and tracking crud in pipe and valves / tracking sources through time
- Identifying fuel failures
- Shielding design and optimization
- Etc.
- Localization of radioactive contamination (nuclear accident, decommissioning)
- Radioactive waste characterization
Geant4 simulation: 60Co and 137Cs hot spots
Nuclear industry
Notes: Compton camera has special advantage of imaging high energy gamma-ray emitting radio-isotope
So it is widely applied in various industries including …

## Slide 6: Application of Compton Camera (2/2)
Homeland security
Large area Compton camera for surveillance
(University of Michigan, 2012)
- Dose verification during in particle therapy
- Nuclear medicine and molecular imaging
Medical applications

## Slide 7: Necessity of “Large-area” Compton Camera
Small Compton camera
Uses small detectors (<10 cm)  low interaction probability  low imaging sensitivity
Large-Area Compton camera
Uses large detectors (a few tens cm)  high interaction probability  high imaging sensitivity (a few tens of times higher)
Source-to-detector distance:
~a few meter
Existing Compton camera size:
<10 cm
Proposed Compton camera size: ~a few tens cm
Notes: There have been some Compton cameras, but most of them are very small.
Small systems have advantage of portability, but basically they suffer from low imaging sensitivity.
As you can easily imagine, the larger system size, the higher imaging sensitivity.

## Slide 8: Enlarging Component Detectors
Example: Dual particle imager (University of Michigan)
Achieved large size (a few tens cm) and cost effectiveness
Low fill factor of each layer which causes loss of sensitivity
Poor spatial resolution (100–140 mm pitch) which causes deterioration of image resolution
No example
Achievable outstanding spatial resolution (a few mm pitch) and large size (a few tens cm)
High production cost for both the pixelated crystal and multiple photo-sensors
Difficult to manufacture and maintain the crystal array structure
Notes: But enlarging component detectors is quite tricky and difficult.
Simple idea is to use of multiple detectors like this, but the fill factor is too low and the spatial resolution is also poor. Whole system size could be enlarged easily, but hard to earn a merit.
Another idea is to use pixelated detector. But this is very hard to be applicable in real world because it needs too high production cost, moreover, it is difficult to manufacture and handling.

## Slide 9: Conceptual design of Large-area Compton camera
27 cm
Large monolithic scintillator with an array of photomultiplier tubes
Range scanner
Large monolithic scintillation detectors crystal: NaI(Tl) (27×27 cm2) photo-sensor: 6×6 PMT array (5.1×5.1 cm2)
Achievable large size (a few tens cm) and cost effectiveness
Achievable spatial resolution as excellent as a pixelated one, by using interaction position estimating algorithm
Effective area: 27×27 cm2
Notes: Here, the conceptual design of the large-area Compton camera is like this.
The large-size monolithic scintillation detector composed of NaI(Tl) crystal.
The size

## Slide 10: Component Detector
Monolithic crystal
  Order-made. Scintitech, MA, USA
  Crystal dimension 27×27×2 cm3 (for scatter) 27×27×3 cm3 (for absorber)
  Material: NaI(Tl)
  Optical window dimension: 30×30×1 cm3
Photomultiplier tube (PMT) array
  PMT: XP3290, Photonis, MA, USA
  Entrance window dimension: □ 51×51 mm2
  Window material: borosilicate glass
  Cathode dimension: □ 48×48 mm2
  Array: □ 6×6

## Slide 11: Interaction Position Estimation in Monolithic Scintillator
How to know where the interaction position was?

by Maximum Likelihood Position Estimation (MLPE) algorithm
-ray
Monolithic crystal
PMT array
Scintillation lights
Interacted position
  Can be estimated by analyzing the dispersion pattern of the scintillation lights

## Slide 12: MLPE Procedure
Pre-process: look-up table generation
 | PMT #1 | PMT #2 | … | PMT #36
Ref. pos. #1 | 0.2467 | 0.2033 | … | 0.0022
Ref. pos. #2 | 0.2433 | 0.2089 | … | 0.0022
Ref. pos. #3 | 0.2384 | 0.2152 | … | 0.0022
… |  |  |  | 
 | PMT #1 | PMT #2 | … | PMT #36
Ref. pos. #1 | 0.2467 | 0.2033 | … | 0.0022
Ref. pos. #2 | 0.2433 | 0.2089 | … | 0.0022
Ref. pos. #3 | 0.2384 | 0.2152 | … | 0.0022
… |  |  |  | 
Measured value on th detector channel for unknown position
 | PMT #1 | PMT #2 | … | PMT #36 | Likelihood
… |  |  |  |  | 
Ref. pos. #27541 | 0.0048 | 0.0113 | … | 0.0035 | 0.016748
Ref. pos. #27542 | 0.0047 | 0.0112 | … | 0.0035 | 0.016750
Ref. pos. #27543 | 0.0046 | 0.0110 | … | 0.0035 | 0.016742
… |  |  |  |  | 
 | PMT #1 | PMT #2 | … | PMT #36 | Likelihood
… |  |  |  |  | 
Ref. pos. #27541 | 0.0048 | 0.0113 | … | 0.0035 | 0.016748
Ref. pos. #27542 | 0.0047 | 0.0112 | … | 0.0035 | 0.016750
Ref. pos. #27543 | 0.0046 | 0.0110 | … | 0.0035 | 0.016742
… |  |  |  |  | 
 | PMT #1 | PMT #2 | … | PMT #36
Measured | 0.0047 | 0.0112 | … | 0.0035
best
estimation
Example)
Estimation process
Look-up table
Virtual grid

## Slide 13: Scintillator Modeling (1/2)
Light yield
Refractive Index
Refractive Index
Transmission
Quantom Efficiency

## Slide 14: Scintillator Modeling (2/2)
*Downscaled to 1/30 light yield in these pictures

## Slide 15: Experimental Results

## Slide 16: Imaging System Optimization
Simulation toolkit: Geant4 (version 10.3)
Physics model: G4EmLivermorePhysics
Radiation source:
  0.662 MeV
  isotropic gamma-ray source
  1 m in front of the system
  108 particles (equivalent to ~10 µCi 137Cs×5 min. measurement)
Thickness of scatter: 0.5, 1.0, 1.5, 2.0, 2.5, 3.0, 3.5, 4.0 cm
Thickness of absorber: 0.5, 1.0, 1.5, 2.0, 2.5, 3.0, 3.5, 4.0 cm
Distance between scatter & absorber: 20 cm, 25 cm, 30 cm

## Slide 17: Performance Parameters
Imaging sensitivity
Resolution FWHM of the point source image
To consider both sensitivity and resolution, we used FOM (figure-of-merit) calculate as
Additionally, considering the price of the detector, FOM per crystal volume is used.

## Slide 18: Optimization Results

## Slide 19: Large-Area Compton Camera
Source
Source
Source

## Slide 20: Gamma-ray Spectrum Generator

## Slide 21: Machine Learning for Gamma Spectroscopy
Many recent studies on gamma spectroscopy have adopted machine learning techniques as increasing in accessibility of the techniques.
To develop gamma spectroscopy algorithms based on machine learning, it requires a considerable amount of gamma-ray spectrum data with known radio-isotope concentrations.

## Slide 22: Data, More Data!

## Slide 23: Experimental Environment
HPGe Detector
  CANBERRA
  Model No.: GR6022
  Dimension: Φ67.3 mm × 65.1 mm
Experimental environment
  5-cm-thick Pb block to shield external radiation
  Point source placed in front of 7.5 cm from the detector head following main axis
  Source: 137Cs 370.3 kBq, 60Co 212.3 kBq

## Slide 24: Monte Carlo Simulation
Monte Carlo Simulation
  Geant4 version 11.0
  Detailed geometry modeling corresponding to the datasheet & experimental environment
  G4EmLivermorePhysics
  1×108 primary particles

## Slide 25: Limitation of Naïve Monte Carlo Simulation
Takes too long time
Do not reflect characteristics of detector
  coincidence sum, dead time loss
  energy resolution

## Slide 26: Spectrum Generation Procedure (1/2)
Save data as list-mode format
  A nuclide of radio-isotope is set to be a primary particle.
  A single decay == an event number
  Record (event #, deposited energy) pair, rather than counts per energy bin in general simulation
Refine data
  Event number  Measured time of the detector (activity, timing resolution)
  Data refining considering charge collection time & dead-time(coincidence sum, count loss)
  Data refining considering energy resolution

## Slide 27: Spectrum Generation Procedure (2/2)
Same process as single nuclide data
Need to sorting process as measured time when combining multiple nuclide data
The final list-mode data can be directly converted to an energy spectrum.

## Slide 28: Time Comparison
A naïve Monte Carlo simulation took about a few hours per a spectrum for 1×108 of primary particles. (i9-11900K @3.50GHz, single core)
The generator takes about 10–20 minutes to generate a spectrum.
  The generator requires single nuclide spectra for every isotope as prerequisites.
  After then, it took 10–20 min. per a spectrum.

## Slide 29: Validation Results
Cs-137 only
Co-60 only
Cs-137 + Co-60
Photopeak counts: 98.8% (662 keV; 137Cs), 98.9% (1173 keV; 60Co), 99.0% (1332 keV; 60Co)
Whole range(0–3000 keV) similarity: 94.3% (137Cs), 94.6% (60Co) (c.f. naïve MC: ~30%)

## Slide 30: Closing

## Slide 31: Closing

## Slide 32
Thank You